## Start Here
Vitalis CMMS — Setup & Configuration Data
Master reference data to load before go-live · Hospital Equipment Maintenance Management System
How to use this workbook
• Each tab is one configuration area. Fill in the columns; the headers show what the system expects.
• Rows shaded BLUE are recommended default values — you can keep them as-is.
• Rows shaded YELLOW are EXAMPLES showing the format — replace them with your hospital's data.
• The 'Configuration Register' tab lists every configurable feature and the order to load it in.
• Load top-to-bottom by phase: foundations first (org, taxonomy), then codes/rules, then templates.
Colour legend
Recommended default — keep
Example — replace with your data
Load phases
Phase 1 — Foundations | Hospitals, buildings, floors, departments, rooms, cost centers, asset taxonomy, manufacturers/models
Phase 2 — Classification | Criticality levels, risk scoring, priorities, statuses, units of measure, numbering sequences
Phase 3 — Codes & Rules | Failure/problem/cause/remedy codes, maintenance & WO types, reason codes, SLA & escalation rules
Phase 4 — Templates | PM templates, calibration & safety-test templates, checklists, custom fields, forms, notifications
Phase 5 — People & Access | Roles & permissions, teams, workshops, skills, certifications, shifts, holidays, maintenance windows
Phase 6 — Commercial | Vendors, contracts, evaluation criteria, stock categories, part classifications, reorder rules, stores

## Configuration Register
Configuration Register
Every configurable feature in the system, grouped by domain, with load order and the role that owns it.
# | Domain | Configurable Item | What it controls | Example values | Depends on | Phase | Owner role
1 | Organization | Hospitals / Branches | Top-level facilities in the group | Cedar Ridge Main, Cedar Ridge East | — | 1 | System Admin
2 | Organization | Buildings / Floors / Rooms | Physical location hierarchy | Tower A > Floor 3 > Room 312 | Hospitals | 1 | CMMS Admin
3 | Organization | Departments / Units | Clinical & support departments | ICU, Radiology, OR, ER, CSSD | Hospitals | 1 | CMMS Admin
4 | Organization | Cost Centers | Financial ownership of assets & costs | CC-ICU-01, CC-RAD-02 | Departments | 1 | Finance
5 | Asset Taxonomy | Asset Categories | Top grouping of equipment | Ventilator, Imaging, Infusion | — | 1 | Biomed Manager
6 | Asset Taxonomy | Subcategories | Finer grouping under category | MRI, CT, Ultrasound (under Imaging) | Categories | 1 | Biomed Manager
7 | Asset Taxonomy | Equipment Groups | Cross-cutting groupings | Life-support, High-voltage | Categories | 1 | Biomed Manager
8 | Asset Taxonomy | Manufacturers | Approved manufacturer list | Medtronic, Philips, GE, Siemens | — | 1 | CMMS Admin
9 | Asset Taxonomy | Models | Model catalog per manufacturer | PB980, MX750, Vida 3T | Manufacturers | 1 | CMMS Admin
10 | Asset Taxonomy | Technical Field Sets | Which technical fields show per category | Ventilator: FiO2, tidal volume… | Categories | 4 | CMMS Admin
11 | Classification | Criticality Levels | Clinical/operational criticality tiers | Life support, High, Medium, Low | — | 2 | Biomed Manager
12 | Classification | Risk Scoring Rules | Factors & weights for risk score | function×5, patient impact×4… | Criticality | 2 | Quality / Risk
13 | Classification | Risk Tiers / Thresholds | Score bands that map to actions | ≥85 Critical, 65-84 High… | Risk rules | 2 | Quality / Risk
14 | Classification | Priorities | Priority levels for requests/WOs | P1 Emergency … P5 Low | — | 2 | CMMS Admin
15 | Statuses | Asset Statuses | Lifecycle states of equipment | Active, Under Maint, Out of Service | — | 2 | CMMS Admin
16 | Statuses | Work-Order Statuses | States a WO moves through | Open, Assigned, Testing, Closed | — | 2 | CMMS Admin
17 | Statuses | Request Statuses | States of a service request | New, Under Review, Converted | — | 2 | CMMS Admin
18 | Codes | Failure Codes | Standardised failure classification | Mechanical, Electrical, Software | — | 3 | Biomed Manager
19 | Codes | Problem Codes | Reported problem categories | No power, Alarm, Inaccurate reading | — | 3 | Biomed Manager
20 | Codes | Cause Codes | Root-cause classification | Worn part, User error, Power surge | — | 3 | Biomed Manager
21 | Codes | Remedy / Action Codes | What was done to fix | Replaced part, Recalibrated | — | 3 | Biomed Manager
22 | Codes | Maintenance Types | Nature of the maintenance | Corrective, Preventive, Calibration | — | 3 | CMMS Admin
23 | Codes | Work-Order Types | Category of work order | Repair, PM, Inspection, Install | Maint. types | 3 | CMMS Admin
24 | Rules | SLA Rules | Response/resolution targets | P1: 15min resp / 4h resolve | Priorities, Criticality | 3 | CMMS Admin
25 | Rules | SLA Pause Reasons | Approved clock-pause reasons | Waiting vendor, Waiting part | — | 3 | CMMS Admin
26 | Rules | Escalation Rules | Who is alerted when & why | Overdue P1 → notify manager | SLA rules | 3 | CMMS Admin
27 | Rules | Approval Workflows & Limits | Who approves what, up to what value | WO > $5,000 → manager approval | Roles | 3 | CMMS Admin
28 | Rules | Reorder Rules | Auto-reorder triggers for parts | Qty ≤ min → raise requisition | Parts | 6 | Storekeeper
29 | Rules | Validation Rules | Field/record validation logic | Closed WO must have labour hrs | Fields | 4 | CMMS Admin
30 | Templates | PM Templates | Reusable PM plans & frequencies | Ventilator Quarterly PM | Categories, Checklists | 4 | Biomed Manager
31 | Templates | Calibration Templates | Calibration procedures & standards | Defib IEC 60601-2-4 | Categories | 4 | Biomed Manager
32 | Templates | Safety-Test Templates | Electrical/performance safety tests | IEC 62353 electrical safety | Categories | 4 | Biomed Manager
33 | Templates | Checklists | Ordered check items with pass/fail | Visual, Functional, Electrical | — | 4 | Biomed Manager
34 | Templates | Inspection Templates | Acceptance & routine inspections | Incoming acceptance checklist | Checklists | 4 | Biomed Manager
35 | Forms | Custom Fields | Extra data fields per record type | Warranty ref, Donation source | Categories | 4 | CMMS Admin
36 | Forms | Mandatory / Conditional Fields | Field requirement rules | Serial # mandatory for imaging | Custom fields | 4 | CMMS Admin
37 | Forms | Form Layouts | Screen layout per role/category | Technician mobile WO form | Custom fields | 4 | CMMS Admin
38 | People | Roles | RBAC role definitions | Biomed Engineer, Storekeeper | — | 5 | System Admin
39 | People | Permissions per Role | Module × action grants | WO: view/create/edit/approve | Roles | 5 | System Admin
40 | People | Data Scopes | Hospital/dept/location restrictions | Tech limited to ICU + ER | Roles, Org | 5 | System Admin
41 | People | Maintenance Teams | Groupings of technicians | Biomedical, Facilities, Imaging | Departments | 5 | Biomed Manager
42 | People | Workshops | Physical maintenance locations | Biomed Workshop L1 | Hospitals | 5 | Biomed Manager
43 | People | Technician Skills | Competency areas | Ventilators, MRI, HVAC | — | 5 | Biomed Manager
44 | People | Certifications | Required certs & expiry tracking | IEC 62353, MRI Safety | Skills | 5 | Biomed Manager
45 | People | Shift Schedules | Working patterns | Day 07-15, Night 15-23 | — | 5 | CMMS Admin
46 | People | Holiday Calendars | Non-working days for SLA | National + hospital holidays | — | 5 | CMMS Admin
47 | People | Maintenance Windows | Allowed servicing windows | OR: after 20:00 only | Departments | 5 | Biomed Manager
48 | Parts | Stock Categories | Grouping of spare parts | Consumable, Repairable, Critical | — | 6 | Storekeeper
49 | Parts | Units of Measure | Issue/stock units | Each, Box, Metre, Litre | — | 2 | Storekeeper
50 | Parts | Part Classifications | ABC / criticality of parts | A-critical, B, C | — | 6 | Storekeeper
51 | Parts | Stores / Warehouses / Bins | Physical stock locations | Central Store > Bin A-12-3 | Hospitals | 6 | Storekeeper
52 | Commercial | Vendors | Service & parts suppliers | Siemens, Medtronic, local ISO | — | 6 | Procurement
53 | Commercial | Contracts | Service contracts & coverage | Full-service MRI 2027 | Vendors, Assets | 6 | Procurement
54 | Commercial | Vendor Evaluation Criteria | How vendors are scored | Response, quality, cost | — | 6 | Quality / Risk
55 | Commercial | Cost Categories | Buckets for cost tracking | Labour, Parts, Vendor, Contract | Cost centers | 6 | Finance
56 | Reasons | Cancellation Reasons | Why a WO/request is cancelled | Duplicate, No fault found | — | 3 | CMMS Admin
57 | Reasons | Deferment Reasons | Why PM is deferred | Equipment in use, No access | — | 3 | CMMS Admin
58 | Reasons | Disposal Reasons | Why equipment is decommissioned | End of life, Beyond repair | — | 3 | CMMS Admin
59 | Reasons | Stock Adjustment Reasons | Why stock is corrected | Damaged, Expired, Miscount | — | 3 | Storekeeper
60 | Reasons | Downtime Reasons | Why equipment is down | Repair, Waiting parts, Vendor | — | 3 | CMMS Admin
61 | System | Notification Templates | Email/SMS/push message content | 'PM due in 7 days' email | Rules | 4 | CMMS Admin
62 | System | Numbering Sequences | Auto-number formats per record | WO-#####, EQ-###### | Hospitals | 2 | CMMS Admin
63 | System | Dashboards & KPI Formulas | Dashboard tiles & KPI definitions | Uptime, PM compliance, MTBF | — | 4 | Biomed Manager
64 | System | Report Templates | Saved report definitions | PM compliance by dept | — | 4 | Biomed Manager
65 | System | Languages | UI languages available | English, Arabic | — | 2 | System Admin
66 | System | Date / Time / Currency | Regional formats | dd-MMM-yyyy, USD, +03:00 | — | 2 | System Admin
67 | System | Terminology Overrides | Hospital-specific naming | 'Asset' → 'Device' | — | 2 | CMMS Admin

## 1. Org & Locations
1. Organization & Locations
Build the location hierarchy top-down. Level values: Hospital, Building, Floor, Department, Unit, Room, Cost Center.
Level | Code | Name | Parent Code | Cost Center | Manager / Contact | Notes
Hospital | H-01 | Cedar Ridge Main Campus | — | — | Dr. R. Aoun | Example — replace with your facility
Building | B-A | Tower A | H-01 | — |  | Example
Floor | F-A3 | Tower A · Level 3 | B-A | — |  | Example
Department | D-ICU | Intensive Care Unit | F-A3 | CC-ICU-01 | Head Nurse | Example
Room | R-ICU-07 | ICU Bed 7 | D-ICU | CC-ICU-01 |  | Example
Cost Center | CC-ICU-01 | ICU Operating Budget | H-01 | — | Finance | Example

## 2. Asset Taxonomy
2. Asset Taxonomy
Categories drive technical fields, default criticality and PM strategy. Fill category > subcategory > group.
Category | Subcategory | Equipment Group | Default Criticality | Default PM Strategy | Technical Field Set
Ventilator | ICU Ventilator | Life-support | Life Support | Quarterly PM + Safety | FiO2, Tidal volume, Pressure, Rate
Imaging | MRI | High-voltage | High | Semi-annual + Vendor | Field strength, Gradient, Coils, Chiller
Imaging | CT | High-voltage | High | Semi-annual | kVp, mAs, Tube, Detector
Infusion | Volumetric Pump | — | High | Semi-annual | Flow rate, Occlusion, Battery
Defibrillator | Manual/AED | Life-support | Life Support | Quarterly + Safety | Energy, Charge time, ECG, Pacing
Facility | Generator | Power | Medium | Monthly load test | kVA, Fuel, Transfer time, Runtime hrs

## 3. Manufacturers & Models
3. Manufacturers & Models
Approved manufacturers and their model catalog. Link each model to a category and default PM/lifetime.
Manufacturer | Model | Category | Typical Life (yrs) | Default PM Template | Notes
Medtronic | PB980 | Ventilator | 10 | Ventilator Quarterly PM | Example
Philips | IntelliVue MX750 | Monitoring | 8 | Monitor Semi-annual PM | Example
Siemens Healthineers | MAGNETOM Vida 3T | Imaging | 12 | MRI Vendor PM | Example
GE HealthCare | Aisys CS2 | Anesthesia | 10 | Anesthesia PM | Example
BD | Alaris 8015 | Infusion | 7 | Infusion Semi-annual PM | Example

## 4. Criticality & Risk
4. Criticality & Risk Scoring
Three tables: criticality levels, risk-scoring factors & weights, and the score bands (tiers).
A · Criticality Levels
Level | Description | Default Priority | Default PM Frequency | Colour
Life Support | Failure endangers life directly | P1 | Quarterly + safety test | Red
High Risk | Major clinical/operational impact | P2 | Semi-annual | Amber
Medium | Important but has backup | P3 | Semi-annual/Annual | Blue
Low | Minimal patient impact | P4 | Annual | Grey
B · Risk Scoring Factors (rate each asset 1-5, multiply by weight, sum, scale to 100)
Factor | Weight | Rating guide (1 low → 5 high)
Clinical function / life-support | 5 | 1 misc · 3 diagnostic · 5 therapeutic/life-support
Patient impact on failure | 4 | 1 none · 5 severe harm
Failure consequence | 4 | 1 negligible · 5 catastrophic
Backup availability (inverse) | 2 | 1 many spares · 5 no backup
Utilization intensity | 2 | 1 rare · 5 continuous
Maintenance complexity | 1 | 1 simple · 5 highly complex
C · Risk Tiers (score band → action)
Score band | Tier | Drives
85 – 100 | Critical / Life-support | Quarterly PM + safety test; P1/P2; mandatory escalation if overdue
65 – 84 | High | Semi-annual PM; P2; priority scheduling
40 – 64 | Medium | Routine PM; standard priority
0 – 39 | Low | Annual/opportunistic PM

## 5. Priorities & SLA
5. Priorities & SLA Rules
Response = time to first action; Resolution = time to return to service. Targets are per priority (and can vary by criticality/contract).
Priority | Label | Example trigger | Response target | Resolution target | Applies to
P1 | Emergency | Life-support equipment down | 15 minutes | 4 hours | Life-support assets
P2 | Critical | Major clinical impact | 1 hour | 24 hours | High-risk assets
P3 | High | Important service affected | 4 hours | 72 hours | Medium assets
P4 | Normal | Standard failure | 8 hours | 5 days | General
P5 | Low | Minor / cosmetic | 1 day | 10 days | Low-risk assets
SLA Pause Reasons (clock stops while active)
Pause reason | Requires approval? | Notes
Waiting for vendor | Yes | External dependency
Waiting for approved part | Yes | Part on order
Equipment inaccessible | No | In clinical use
Awaiting user/department | No | Access or info needed

## 6. Statuses
6. Statuses
Configurable state lists. 'Type' = which record the status belongs to. Order defines display sequence.
Type | Status | Order | Category | Notes
Asset Status | Pending Acceptance | 1 | Lifecycle
Asset Status | Active | 2 | Lifecycle
Asset Status | Available | 3 | Lifecycle
Asset Status | In Use | 4 | Lifecycle
Asset Status | Under Maintenance | 5 | Lifecycle
Asset Status | Under PM | 6 | Lifecycle
Asset Status | Under Calibration | 7 | Lifecycle
Asset Status | Awaiting Spare Part | 8 | Lifecycle
Asset Status | Awaiting Vendor | 9 | Lifecycle
Asset Status | Out of Service | 10 | Lifecycle
Asset Status | Quarantined | 11 | Lifecycle
Asset Status | Loaned | 12 | Lifecycle
Asset Status | In Storage | 13 | Lifecycle
Asset Status | Transferred | 14 | Lifecycle
Asset Status | Decommissioned | 15 | Lifecycle
Asset Status | Disposed | 16 | Lifecycle
Work-Order Status | Draft | 1 | Corrective/PM
Work-Order Status | Submitted | 2 | Corrective/PM
Work-Order Status | Open | 3 | Corrective/PM
Work-Order Status | Triaged | 4 | Corrective/PM
Work-Order Status | Assigned | 5 | Corrective/PM
Work-Order Status | Accepted | 6 | Corrective/PM
Work-Order Status | In Progress | 7 | Corrective/PM
Work-Order Status | Waiting Parts | 8 | Corrective/PM
Work-Order Status | Waiting Vendor | 9 | Corrective/PM
Work-Order Status | On Hold | 10 | Corrective/PM
Work-Order Status | Testing | 11 | Corrective/PM
Work-Order Status | Verification | 12 | Corrective/PM
Work-Order Status | Closed | 13 | Corrective/PM
Work-Order Status | Cancelled | 14 | Corrective/PM
Request Status | New | 1 | Service request
Request Status | Under Review | 2 | Service request
Request Status | Approved | 3 | Service request
Request Status | Converted to WO | 4 | Service request
Request Status | Rejected | 5 | Service request

## 7. Codes
7. Classification Codes
Failure / problem / cause / remedy codes plus maintenance & work-order types. Used for analytics (repeat failures, MTBF).
Code Type | Code | Description | Applies to category
Failure Code | MECH | Mechanical failure | All
Failure Code | ELEC | Electrical failure | All
Failure Code | SW | Software / firmware | All
Failure Code | SENS | Sensor / accuracy | All
Failure Code | WEAR | Wear & tear | All
Problem Code | NOPWR | No power / won't start | All
Problem Code | ALARM | Persistent alarm | All
Problem Code | INACC | Inaccurate reading | All
Problem Code | LEAK | Leak / fluid | All
Problem Code | NOISE | Abnormal noise | All
Cause Code | WORN | Worn component | All
Cause Code | USER | User error | All
Cause Code | SURGE | Power surge | All
Cause Code | CONTAM | Contamination | All
Cause Code | AGE | End of life | All
Remedy Code | REPL | Replaced part | All
Remedy Code | CAL | Recalibrated | All
Remedy Code | ADJ | Adjusted / tuned | All
Remedy Code | CLEAN | Cleaned | All
Remedy Code | UPD | Updated firmware | All
Maintenance Type | COR | Corrective | All
Maintenance Type | PRE | Preventive | All
Maintenance Type | CAL | Calibration | All
Maintenance Type | SAF | Safety inspection | All
Maintenance Type | EMG | Emergency | All
Work-Order Type | REP | Repair | All
Work-Order Type | PM | Preventive Maintenance | All
Work-Order Type | INS | Inspection | All
Work-Order Type | INST | Installation | All
Work-Order Type | DEC | Decommission | All

## 8. PM & Calibration
8. PM, Calibration & Safety-Test Templates
Reusable maintenance plans. Attach to a category or model. Link a checklist template (tab 9).
Template Type | Name | Applies to (category/model) | Frequency | Standard / Basis | Checklist | Est. duration
PM | Ventilator Quarterly PM | Ventilator | Quarterly | OEM + IEC 62353 | CL-VENT | 90 min
PM | Monitor Semi-annual PM | Monitoring | Semi-annual | OEM | CL-GEN | 45 min
Calibration | Defibrillator Energy Cal | Defibrillator | Annual | IEC 60601-2-4 | CL-DEFIB | 60 min
Calibration | Infusion Pump Flow Cal | Infusion | Annual | IEC 60601-2-24 | CL-INF | 45 min
Safety Test | Electrical Safety Test | All electro-medical | Annual | IEC 62353 | CL-ELEC | 30 min

## 9. Checklists & Fields
9. Checklists, Custom Fields & Forms
Checklist items (one per row) and any extra custom fields. Item type: Check, Reading (with range), or Text.
A · Checklist Items
Checklist ID | Section | Item | Type | Unit | Expected / Range
CL-VENT | Visual | Housing & accessories intact | Check |  | Pass
CL-VENT | Functional | Delivered tidal volume | Reading | mL | 450–550
CL-VENT | Functional | FiO2 accuracy | Reading | % | 19–23
CL-ELEC | Electrical | Earth leakage current | Reading | mA | 0–0.5
CL-DEFIB | Energy | Delivered energy @ 200J | Reading | J | 190–210
B · Custom Fields
Field name | Applies to | Data type | Mandatory? | Condition
Donation Source | Asset | Text | No | If funding = Donation
UDI / GS1 | Asset | Text | Conditional | If category = medical device
Radiation Class | Asset | List | Conditional | If category = Imaging

## 10. Parts Setup
10. Spare Parts Setup
Stock categories, units, classifications, reorder rules and store locations.
A · Stock Categories & Units
Type | Value | Description
Stock Category | Consumable | Single-use, replenished
Stock Category | Repairable | Rotable / repaired & returned
Stock Category | Critical Spare | Must-hold for life-support
Unit of Measure | Each | Individual item
Unit of Measure | Box | Multi-pack
Unit of Measure | Metre | Length
Part Class | A | Critical / high-value — tight control
Part Class | B | Moderate
Part Class | C | Low-value / bulk
B · Stores / Warehouses / Bins
Store | Warehouse | Bin | Hospital | Notes
Central Store | Main WH | A-12-3 | H-01 | Example — replace
Biomed Workshop | Workshop WH | D-01-4 | H-01 | Example

## 11. Teams, Skills, Shifts
11. Teams, Skills, Certifications & Shifts
People-side configuration. Skills gate assignment; certifications track expiry.
A · Maintenance Teams & Workshops
Type | Name | Covers | Location
Team | Biomedical | Medical & lab equipment | H-01
Team | Facilities | HVAC, power, gases | H-01
Team | Imaging | MRI/CT/X-ray | H-01
Workshop | Biomed Workshop L1 | Bench repairs | Tower A L1
B · Skills, Certifications, Shifts, Holidays
Type | Value | Detail
Skill | Ventilators | Competency area
Skill | MRI | Requires MRI Safety cert
Skill | HVAC | Facilities
Certification | IEC 62353 Electrical Safety | Renew every 3 yrs
Certification | MRI Safety Officer | Renew annually
Shift | Day 07:00–15:00 | Weekdays
Shift | Night 23:00–07:00 | On-call rota
Holiday | New Year 01 Jan | Non-working — SLA pauses
Maintenance Window | OR after 20:00 | OR servicing only after hours

## 12. Roles & Access
12. Roles & Data Scopes
Define roles and their scope. Permissions (module × action) are set inside the app's Roles screen.
Role | Scope type | Description | # Users (est.)
System Administrator | Global | Technical config, users, integrations | 2
Hospital / CMMS Administrator | Hospital | Config, lookups, workflows, master data | 3
Clinical Engineering Manager | Hospital | Program, strategy, approvals | 2
Biomedical Engineer | Assigned | Diagnose, maintain, calibrate | 9
Biomedical Technician | Assigned | Execute tasks & checklists | 14
Facility Engineer / Technician | Facilities | HVAC, power, gases | 6
Clinical Department User | Department | Report faults, track requests | 220
Department Manager / Head Nurse | Department | Approve requests, review assets | 34
Storekeeper / Spare Parts Officer | Store | Stock, issues, reorder | 4
Procurement Officer | Hospital | Requisitions, POs, vendors | 3
Finance User | Hospital | Costs, capitalization, budgets | 2
Quality / Patient Safety / Risk | Hospital | Incidents, recalls, compliance | 3
Vendor / External Service Provider | Assigned WO | Restricted portal, assigned work | 11
Auditor / Accreditation | Read-only | Read-only records & certificates | 2
Management / Executive | Hospital | Read-only dashboards & KPIs | 8

## 13. Reasons & Numbering
13. Reason Codes, Notifications & Numbering
Reason lists (audited), notification templates and auto-numbering formats.
A · Reason Codes
Reason Type | Reason | Notes
Cancellation | Duplicate request | Merge with original
Cancellation | No fault found | Closed after check
Deferment | Equipment in clinical use | Reschedule PM
Deferment | No access to location | OR/ICU busy
Disposal | End of useful life | Beyond economic repair
Disposal | Irreparable failure | Safety risk
Stock Adjustment | Damaged in store | Requires supervisor
Downtime | Waiting for parts | Counts to downtime
B · Notification Templates & Numbering
Type | Name / Format | Channel / Trigger | Recipients
Notification | PM Due Reminder | Email · 7 days before due | Assigned team
Notification | SLA Breach Alert | Push+SMS · at 80% of SLA | Manager
Notification | Calibration Expiry | Email · 30 days before | Biomed
Numbering | WO-##### (WO-00001) | Per hospital, resets yearly | —
Numbering | EQ-###### (EQ-100001) | Global asset ID | —
Numbering | SR-##### / PM-##### | Requests & PM WOs | —

## 14. System Settings
14. System Settings
Regional and global defaults.
Setting | Value | Notes
Primary language | English | Add Arabic as secondary if needed
Date format | dd-MMM-yyyy | e.g. 28-Aug-2026
Time zone | (GMT+03:00) Beirut | Used for SLA clocks
Currency | USD ($) | For costs & valuation
Working week | Mon–Fri (+Sat half) | Drives SLA business hours
Fiscal year start | January | Cost reporting
Terminology override | 'Asset' → 'Device' | Optional hospital wording