## Equipment
Equipment
Fill in one row per piece of equipment. Columns marked * are required.
ID*: Unique identifier (e.g. EQ-100001). Tag*: Hospital asset tag. Name*: Equipment name.
Status: inuse, available, maint, retired. Crit: life, high, med, low.
IC (inspection category): asset, defib, pump, ventilator, imaging, etc.
Dept: Department name (must match Departments sheet). SLA: P1, P2, P3, P4, P5.
Warranty: Active or Expired. Track_downtime: TRUE or FALSE.
Ownership_type: Owned, Leased, Rented, Donated. Depreciation_method: straight-line.
Leave dates as YYYY-MM-DD. Leave blank if unknown — defaults will be applied.
id* | tag* | name* | model | mfr | cat | ic | dept | loc | status | crit | risk | pm_interval_days | next_pm | warranty | warranty_exp | cal_due | age | cost | serial | sla | track_downtime | downtime_limit_hours | ownership_type | depreciation_years | depreciation_method | acquisition_date | salvage_value
EQ-100001 | VENT-ICU-01 | Puritan Bennett 980 Ventilator | PB980 | Medtronic | Ventilator | ventilator | ICU | ICU · Bed 4 | inuse | life | 96 | 90 | 2026-12-01 | Active | 2027-06-30 | 2026-09-15 | 3 | 32000 | MDT-PB980-00123 | P1 | TRUE | 24 | Owned | 7 | straight-line | 2023-01-15 | 2000

## Departments
Departments
List all hospital departments. The dept name in the Equipment sheet must match the name here.
ID*: Unique identifier (e.g. dept-icu). Name*: Department name.
id* | name* | description
dept-icu | ICU | Intensive Care Unit
dept-nicu | NICU | Neonatal Intensive Care Unit
dept-or | Operating Room | Operating Theatre
dept-emergency | Emergency | Emergency Department
dept-radiology | Radiology | Radiology and Imaging
dept-nephrology | Nephrology | Nephrology / Dialysis
dept-facilities | Facilities | Facilities Management

## Teams
Teams
Maintenance teams that work orders and PMs can be assigned to.
ID*: Unique identifier. Name*: Team name. Color: CSS color variable or hex.
id* | name* | description | color | sort_order
biomedical | Biomedical | Biomedical engineering team | var(--primary) | 1
imaging | Imaging | Imaging systems team | var(--info) | 2
facilities | Facilities | Facilities maintenance team | var(--ok) | 3
vendor | Vendor | External vendor / contractor | var(--warn) | 4

## Technicians
Technicians
Biomedical / clinical engineering staff who perform maintenance.
ID*: Unique identifier (e.g. U-002). Name*: Full name. Trade*: Biomedical, Imaging, Facilities, etc.
Skills: Comma-separated list (e.g. Ventilators, Defibrillators, Patient Monitors).
Certs: Format as 'Cert Name:expiry-date; Cert Name2:expiry-date' (e.g. IEC 62353:2027-05-01; MRI Safety:2026-09-20).
Load: Current workload (number). Cap: Max capacity (number). Avail: On shift, On call, Off.
id* | name* | trade | skills | certs | load | cap | avail
U-002 | Karim Haddad | Biomedical | Ventilators, Defibrillators, Patient Monitors, Infusion Pumps | IEC 62353 Electrical Safety:2027-05-01; Ventilator OEM — Medtronic:2026-11-15 | 6 | 8 | On shift
U-003 | Maya Saab | Imaging | MRI, CT, Ultrasound, X-Ray | MRI Safety Officer:2026-09-20; Radiation Protection:2028-01-10 | 5 | 8 | On shift
U-004 | Nadia Fares | Biomedical | Patient Monitors, Sterilizers, Lab Equipment | IEC 62353 Electrical Safety:2026-10-05; Sterilizer Validation:2027-03-01 | 4 | 8 | On shift
U-005 | Rami Nassar | Facilities | HVAC, Generators, UPS, Chillers, Medical Gases | Electrical Licence:2027-07-01; Medical Gas Systems:2026-12-01 | 3 | 8 | On call

## PM Checklist Templates
PM Checklist Templates
Each row defines one checklist template that can be linked to PM plans and PM work orders.
ID*: Unique identifier (e.g. pmt-vent-01). Name*: Descriptive name. Description: Optional.
Sections*: A JSON array defining the checklist. See the example below and the 'Checklist Format Guide' sheet.
EXAMPLE sections JSON:
[{"title":"Visual Inspection","items":[{"t":"No visible damage","type":"check"},{"t":"Cables intact","type":"check"}]},{"title":"Electrical Safety","items":[{"t":"Earth leakage","type":"reading","unit":"mA","nominal":0.12,"min":0,"max":0.5}]}]
Item types: 'check' = pass/fail/N-A button. 'reading' = numeric input with pass/fail range.
For 'reading' items: t=task name, unit=unit symbol, nominal=expected value, min/max=acceptable range.
For 'check' items: t=task name only. Unit, nominal, min, max are not needed.
id* | name* | description | sections (JSON)*
pmt-ventilator-01 | Ventilator Quarterly PM | Quarterly preventive maintenance checklist for ICU ventilators | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage to casing", "type": "check" }, { "t": "Cables and connectors intact", "type": "check" }, { "t": "Labels and warnings legible", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage current", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage current", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 }, { "t": "Patient leakage current", "type": "reading", "unit": "\u00b5A", "nominal": 10, "min": 0, "max": 50 } ] }, { "title": "Functional Test", "items": [ { "t": "Power-on self-test passes", "type": "check" }, { "t": "Alarms audible and visible", "type": "check" }, { "t": "Battery backup operational", "type": "check" }, { "t": "Controls and display operate correctly", "type": "check" } ] } ]

## PM Plans
PM Plans
Each row defines a recurring PM schedule linked to a piece of equipment and a checklist template.
ID*: Unique identifier (e.g. PMP-0001). Name*: Plan name. EQ_ID*: Must match an equipment ID from the Equipment sheet.
TPL*: Must match a template ID from the PM Checklist Templates sheet (or a built-in: generic, ventilator, defib, anesthesia, imaging).
Freq*: Monthly, Quarterly, Semi-annual, Annual, or Biennial.
Technician: Full name of the assigned technician (must match Technicians sheet).
Team: Must match a team name from the Teams sheet. Start_date: YYYY-MM-DD.
Active: TRUE or FALSE. Reminder_days: Days before due date to send reminder (e.g. 7).
id* | name* | eq_id* | tpl* | freq* | technician | team | start_date | active | reminder_days
PMP-0001 | Ventilator Quarterly PM | EQ-100001 | pmt-ventilator-01 | Quarterly | Karim Haddad | Biomedical | 2026-09-01 | TRUE | 7
PMP-0002 | Defibrillator Quarterly Safety | EQ-100002 | defib | Quarterly | Karim Haddad | Biomedical | 2026-09-01 | TRUE | 7
PMP-0003 | MRI Semi-annual PM | EQ-100003 | imaging | Semi-annual | Maya Saab | Imaging | 2026-09-01 | TRUE | 7
PMP-0004 | Generator Monthly Load Test | EQ-100004 | generic | Monthly | Rami Nassar | Facilities | 2026-09-01 | TRUE | 7

## PM Work Orders
PM Work Orders
Individual PM jobs generated from PM plans (or created manually). Each links to equipment and a checklist template.
ID*: Unique identifier (e.g. PM-6001). EQ_ID*: Must match equipment ID. Title*: PM job title.
Due*: Due date (YYYY-MM-DD). Freq*: Monthly, Quarterly, Semi-annual, Annual, Biennial.
TPL*: Template ID from PM Checklist Templates sheet or built-in (generic, ventilator, defib, anesthesia, imaging).
Status: scheduled, inprogress, completed, overdue. Team: Must match Teams sheet. Technician: Name from Technicians sheet.
id* | eq_id* | title* | due* | freq* | tpl* | status | team | technician
PM-6001 | EQ-100001 | Ventilator Quarterly PM | 2026-09-15 | Quarterly | pmt-ventilator-01 | scheduled | Biomedical | Karim Haddad
PM-6002 | EQ-100002 | Defibrillator Safety Test | 2026-09-20 | Quarterly | defib | scheduled | Biomedical | Karim Haddad
PM-6003 | EQ-100003 | MRI Semi-annual PM | 2026-10-01 | Semi-annual | imaging | scheduled | Imaging | Maya Saab

## Priorities (Reference)
Priorities — Reference Only
These are the SLA priority levels already configured in the system. Do not change unless you know what you are doing.
Use the Priority column (P1-P5) in the Equipment sheet's SLA column.
priority | label | example_trigger | response_target | resolution_target | resolution_hours | warning_pct | applies_to | color | sort_order
P1 | Emergency | Life-support equipment down | 15 minutes | 4 hours | 4 | 75 | Life-support assets | var(--crit) | 1
P2 | Critical | Major clinical impact | 1 hour | 24 hours | 24 | 75 | High-risk assets | var(--warn) | 2
P3 | High | Important service affected | 4 hours | 72 hours | 72 | 75 | Medium assets | var(--info) | 3
P4 | Normal | Standard failure | 8 hours | 5 days | 120 | 75 | General | var(--text-3) | 4
P5 | Low | Minor / cosmetic | 1 day | 10 days | 240 | 75 | Low-risk assets | var(--text-3) | 5

## Criticality (Reference)
Criticality Levels — Reference Only
These are the criticality levels used in the Equipment sheet's Crit column.
Use the ID column value (life, high, med, low) in the Equipment sheet.
id | level | description | default_priority | default_pm_frequency | risk_score | sort_order
life | Life Support | Failure endangers life directly | P1 | Quarterly + safety test | 90 | 1
high | High Risk | Major clinical/operational impact | P2 | Semi-annual | 75 | 2
med | Medium | Important but has backup | P3 | Semi-annual/Annual | 50 | 3
low | Low | Minimal patient impact | P4 | Annual | 50 | 4

## PM Frequencies (Reference)
PM Frequencies — Reference Only
Available PM frequency options. Use the Label value in the PM Plans and PM Work Orders sheets.
id | label | months_interval | sort_order
monthly | Monthly | 1 | 1
quarterly | Quarterly | 3 | 2
semiannual | Semi-annual | 6 | 3
annual | Annual | 12 | 4
biennial | Biennial | 24 | 5

## Units of Measure (Reference)
Units of Measure — Reference Only
Available units for checklist 'reading' type items. Use the Symbol value in the unit field of checklist sections.
id | name | symbol | value_type | sort_order
ohm | Ohms | Ω | number | 1
milliamp | Milliamps | mA | number | 2
microamp | Microamps | µA | number | 3
milliliter | Milliliters | mL | number | 4
cmh2o | cmH₂O | cmH₂O | number | 5
percent | Percent | % | number | 6
bpm | Breaths per min | bpm | number | 7
joules | Joules | J | number | 8
seconds | Seconds | s | number | 9
kvp | kVp | kVp | number | 10
mgy | Milligray | mGy | number | 11
tempc | Temperature (°C) | °C | number | 12
tempf | Temperature (°F) | °F | number | 13
pressure_bar | Bar | bar | number | 14
pressure_psi | PSI | psi | number | 15
voltage | Volts | V | number | 16
frequency | Hertz | Hz | number | 17
rpm | RPM | RPM | number | 18
boolean | Pass/Fail | — | boolean | 19
text | Text/Note | — | text | 20

## Built-in Templates (Reference)
Built-in Checklist Templates — Reference Only
These templates are already built into the system. You can reference them by their key in the PM Plans/PM Work Orders 'tpl' column.
You can also create custom templates in the PM Checklist Templates sheet.
template_key | name | description | sections (JSON)
generic | Generic PM | Visual + electrical safety + functional test for any equipment | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage", "type": "check" }, { "t": "Cables intact", "type": "check" }, { "t": "Labels legible", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 } ] }, { "title": "Functional Test", "items": [ { "t": "Power-on self-test passes", "type": "check" }, { "t": "Alarms audible & visible", "type": "check" }, { "t": "Battery backup operational", "type": "check" }, { "t": "Controls & display operate correctly", "type": "check" } ] } ]
ventilator | Ventilator PM | For ICU ventilators — includes tidal volume, airway pressure, FiO2 readings | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage", "type": "check" }, { "t": "Cables intact", "type": "check" }, { "t": "Labels legible", "type": "check" } ] }, { "title": "Functional Performance", "items": [ { "t": "Delivered tidal volume", "type": "reading", "unit": "mL", "nominal": 500, "min": 450, "max": 550 }, { "t": "Airway pressure", "type": "reading", "unit": "cmH\u2082O", "nominal": 20, "min": 18, "max": 22 }, { "t": "FiO\u2082 accuracy", "type": "reading", "unit": "%", "nominal": 21, "min": 19, "max": 23 }, { "t": "Respiratory rate", "type": "reading", "unit": "bpm", "nominal": 12, "min": 11, "max": 13 }, { "t": "Apnoea & disconnect alarms operate", "type": "check" }, { "t": "O\u2082 cell calibrated", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 } ] } ]
defib | Defibrillator PM | For defibrillators — includes energy delivery, charge time, ECG, pacing checks | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage", "type": "check" }, { "t": "Cables intact", "type": "check" }, { "t": "Labels legible", "type": "check" } ] }, { "title": "Energy & Function", "items": [ { "t": "Delivered energy @ 200 J", "type": "reading", "unit": "J", "nominal": 200, "min": 190, "max": 210 }, { "t": "Charge time to 200 J", "type": "reading", "unit": "s", "nominal": 7, "min": 0, "max": 10 }, { "t": "ECG display accurate & clear", "type": "check" }, { "t": "Pacing function operational", "type": "check" }, { "t": "Battery fully charged & holds", "type": "check" }, { "t": "Self-test log reviewed", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 } ] } ]
anesthesia | Anesthesia Machine PM | For anesthesia machines — gas concentration, vaporizer, circuit leak checks | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage", "type": "check" }, { "t": "Cables intact", "type": "check" }, { "t": "Labels legible", "type": "check" } ] }, { "title": "Gas & Function", "items": [ { "t": "O\u2082 concentration delivered", "type": "reading", "unit": "%", "nominal": 100, "min": 95, "max": 100 }, { "t": "Vaporizer output", "type": "reading", "unit": "%", "nominal": 2, "min": 1.8, "max": 2.2 }, { "t": "Circuit leak < 200 mL/min", "type": "check" }, { "t": "Scavenging system functional", "type": "check" }, { "t": "O\u2082 failure alarm operates", "type": "check" }, { "t": "Ventilator mode functional", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 } ] } ]
imaging | Imaging Equipment PM | For MRI, CT, X-ray, ultrasound — image quality, kVp, radiation output | [ { "title": "Visual Inspection", "items": [ { "t": "No visible damage", "type": "check" }, { "t": "Cables intact", "type": "check" }, { "t": "Labels legible", "type": "check" } ] }, { "title": "Performance & Radiation Safety", "items": [ { "t": "Image-quality phantom test passes", "type": "check" }, { "t": "kVp accuracy", "type": "reading", "unit": "kVp", "nominal": 80, "min": 76, "max": 84 }, { "t": "Radiation output", "type": "reading", "unit": "mGy", "nominal": 5, "min": 4.5, "max": 5.5 }, { "t": "Emergency stop functional", "type": "check" }, { "t": "Interlocks & warning lights operational", "type": "check" } ] }, { "title": "Electrical Safety", "items": [ { "t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5 }, { "t": "Enclosure leakage", "type": "reading", "unit": "mA", "nominal": 0.1, "min": 0, "max": 0.5 } ] } ]

## Checklist Format Guide
Checklist Sections JSON Format Guide
Overview
The 'sections' field in PM Checklist Templates is a JSON array of section objects.
Each section has a 'title' and an 'items' array.
Each item has a 't' (task name) and a 'type' ('check' or 'reading').
'check' type item
Used for pass/fail/N-A tasks (no numeric measurement needed).
Fields: t = task description, type = 'check'
Example: {"t": "Power-on self-test passes", "type": "check"}
'reading' type item
Used for numeric measurements with an acceptable range.
Fields: t = task name, type = 'reading', unit = unit symbol (see Units of Measure sheet),
nominal = expected/ideal value, min = minimum acceptable, max = maximum acceptable.
Example: {"t": "Delivered tidal volume", "type": "reading", "unit": "mL", "nominal": 500, "min": 450, "max": 550}
Full example
[
{
"title": "Visual Inspection",
"items": [
{"t": "No visible damage", "type": "check"},
{"t": "Cables intact", "type": "check"}
]
},
{
"title": "Electrical Safety",
"items": [
{"t": "Earth leakage", "type": "reading", "unit": "mA", "nominal": 0.12, "min": 0, "max": 0.5}
]
}
]
Available units
mA (milliamps), µA (microamps), mL (milliliters), cmH₂O, %, bpm, J (joules),
s (seconds), kVp, mGy (milligray), °C, °F, bar, psi, V (volts), Hz, RPM, Ω (ohms)
Use '—' for unitless check items.

## Spare Parts
Spare Parts Inventory
Fill in one row per spare part in your inventory.
ID*: Unique part number (e.g. P-0001). Name*: Part name. Mfr: Manufacturer.
Cat: Electrical, Mechanical, Pneumatic, Electronic, Consumable, Sensor, Filter, Other.
Qty: Current stock. Min_qty: Reorder threshold. Max_qty: Maximum stock level.
Bin: Storage location. Cost: Unit cost. Crit: TRUE for critical spares (life-support).
SAP_PO_number: Optional SAP purchase order reference.
id* | name* | mfr | cat | qty | min_qty | max_qty | bin | cost | crit | sap_po_number
P-0001 | Infusion Pump Battery Module | BD | Electronic | 5 | 2 | 10 | A-12 | 145 | TRUE | SAP-PO-2024-0150
P-0002 | Ventilator Bacterial Filter | Medtronic | Consumable | 20 | 5 | 50 | B-04 | 18 | FALSE | SAP-PO-2024-0151
P-0003 | Patient Monitor ECG Cable | Philips | Electrical | 8 | 3 | 15 | C-07 | 65 | FALSE

## Vendors
Vendors / Contractors
External service providers and contractors.
ID*: Unique identifier. Name*: Vendor name. Cat: OEM, ISO, Contractor, Supplier.
Contract: Contract reference. SLA: Compliance percentage. Open: Open work orders count.
Cost: Annual cost. Exp: Contract expiry date (YYYY-MM-DD).
id* | name* | cat | contract | sla | open | cost | exp
V-001 | Medtronic Service | OEM | MSC-2026-001 | 98 | 2 | 45000 | 2027-01-31
V-002 | Siemens Healthineers | OEM | SHS-2026-003 | 95 | 1 | 120000 | 2026-12-31
V-003 | BioMed Engineering Services | ISO | BES-2026-002 | 92 | 0 | 28000 | 2027-06-30

## Asset Categories
Asset Categories
Category definitions for equipment classification. Used to apply defaults for criticality and PM strategy.
ID*: Leave blank for new entries (auto-generated UUID). Category*: e.g. Ventilator, Infusion, Imaging.
Subcategory: e.g. ICU Ventilator, Volumetric Pump, MRI. Equipment_group: e.g. Life-support, High-voltage.
Default_criticality: life, high, med, low. Default_pm_strategy: e.g. Quarterly PM + Safety.
id | category* | subcategory | equipment_group | default_criticality | default_pm_strategy
| Ventilator | ICU Ventilator | Life-support | life | Quarterly PM + Safety
| Defibrillator | Manual/AED | Life-support | life | Quarterly + Safety
| Infusion | Volumetric Pump | — | high | Semi-annual
| Imaging | MRI | High-voltage | high | Semi-annual + Vendor
| Imaging | CT | High-voltage | high | Semi-annual
| Facility | Generator | Power | med | Monthly load test

## Competencies (Reference)
Competencies — Reference Only
Skills/competencies that can be assigned to work order types and technicians.
id | name | sort_order
ventilators | Ventilators | 1
defibrillators | Defibrillators | 2
patient_monitors | Patient Monitors | 3
infusion_pumps | Infusion Pumps | 4
mri | MRI | 5
ct | CT | 6
ultrasound | Ultrasound | 7
x-ray | X-Ray | 8
sterilizers | Sterilizers | 9
hvac | HVAC | 10
generators | Generators | 11
medical_gases | Medical Gases | 12
dialysismachines | Dialysis machines | 13

## WO Types (Reference)
Work Order Types — Reference Only
Types of work orders that can be created. Each type lists required competencies.
id | name | competencies | sort_order
corrective | Corrective | Ventilators, Defibrillators, Patient Monitors, Infusion Pumps, Sterilizers, Medical Gases, Dialysis machines | 1
preventive | Preventive | Ventilators, Defibrillators, Patient Monitors, Infusion Pumps, Sterilizers, Medical Gases, MRI, CT, Ultrasound, X-Ray, HVAC, Generators, Dialysis machines | 2
calibration | Calibration | Patient Monitors, Infusion Pumps, MRI, CT, Ultrasound, X-Ray, Dialysis machines | 3
safety_test | Safety Test | Defibrillators, Patient Monitors, Sterilizers, Medical Gases | 4

## Users
Users
System users who can log in. Each user needs a role (see Roles sheet).
ID*: Unique identifier (e.g. U-101). Name*: Full name. Email*: Login email.
Role*: Must match a role name from the Roles sheet (e.g. Superadmin, Biomedical Supervisor, Technician).
Dept: Department name (must match Departments sheet). Status: active or inactive.
Note: Passwords are set through the login system, not in this sheet.
id* | name* | email* | role* | dept | status
U-101 | Iman Katergi | imankatergi@gmail.com | Biomedical Supervisor | North one | active
U-102 | Cherif Katerji | cherif@makassedhospital.org | Biomedical Engineer | North one | active
U-103 | Karim Haddad | karim@makassedhospital.org | Technician | ICU | active

## Roles
Roles
User roles that define permissions. Each role maps to a set of module/action permissions.
ID*: Unique identifier. Name*: Role name. Description: What this role can do.
Scope: hospital or department. System: TRUE for built-in system roles.
Email notification preferences: TRUE/FALSE for each event type.
id* | name* | description | scope | system | email_sr_create | email_sr_update | email_sr_close | email_wo_create | email_wo_update | email_wo_close | email_pm_create | email_pm_update | email_pm_close
role-superadmin | Superadmin | Full system access | hospital | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
role-bio-supervisor | Biomedical Supervisor | Supervises biomedical team, manages SRs and WOs | hospital | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE
role-technician | Technician | Performs maintenance work | hospital | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE

## Workflows (Reference)
Workflows — Reference Only
Predefined work order state machines. Work orders follow these step sequences.
States: JSON array of workflow states. Step_config: JSON for step-level configuration.
id | name | states (JSON) | step_config (JSON)
wf-1 | Standard Corrective | [ "triaged", "assigned", "accepted", "inprogress", "pending_closeout", "closed" ] | {}